## Slide 1: Gaseous detectors: from past to present
V. Peskov
CERN

## Slide 2: Topics to be covered:
History of detectors of photons and charged particles
Main designs of gaseous detectors
Characteristics and physics of operation
Applications

## Slide 3: Modern detectors of photons and charged particles can be classified into four main categories: vacuum, gaseous, liquid and solid–state.
Introduction

## Slide 4: Images of the most advanced detectors

## Slide 5
Vacuum amplifier
Gas electron amplifier
Avalanches in liquid
Solid-state amplifier

## Slide 6: There is no universal detector which is the best for all applications
For each specific application exist one or several optimal detectors

## Slide 7
Generic operational schematics of electronic detectors of photons and charged particles
(Removable)
(…and drift of primary ions)
(in many new designs!)
However , most of the detectors have a common operational schematics

## Slide 8: F
Example: solid-state avalanche photodiode

## Slide 9: Why do we discuss gaseous detectors today ? What is the interest associate to them?
From historical point of view: they were the first electronic detectors
From theoretical point of view : after learning you will be easy to understand solid –state and liquid detectors
From practical point of view: large area

## Slide 10: Gaseous detectors of photons and charged particles due to their features such as high gas gains, good position and energy resolutions and low cost per unit of sensitive area occupy a special niche among other detectors

## Slide 11: To appreciate these advantages it will be useful to consider the history of developing detectors of photon and charge particles

## Slide 12: History of detectors of charged particles and photons
Part-1

## Slide 13: Photographic films were the first detectors used in studies of X-rays by Roentgen (1895)and radioactivity by Becquerel (1986)
Becquerel had discovered in 1896 that uranium gives off invisible rays that fog photographic film. The Curies had shown that thorium acts similarly. In November 1898, they reported the chemical separation of two unknown materials from the ore pitchblende, both highly radioactive — the first use of the word.

## Slide 14: At this time the photographic technique did not allow to detect a single photon or a single particle

## Slide 15: Around 1910 Rutherford start using ZnS scintillator
allowing to detect single alpha particles
Zinc sulphide fluoresces (gives out a photon of visible light) when it is hit by a charged particle. Covering the microscope lens with ZnS allowed the viewer to 'see' where the alpha particles hit (or at least count their impacts).

## Slide 16: Scintillation screens

## Slide 17: First electronic detector of photons and charged particles-Geiger counter

## Slide 18: What was already known at this time and what physical principles were exploited in the Geiger counter?

## Slide 19
Note:
Electrons drift much faster than ions
Ions movement in electric fields was already observed and understood.
Ionization chambers were already used in some experiments

## Slide 20: Note that ionization chamber were later and still now used in medicine as dosimeters
Ionization chamber current vs. the applied voltage

## Slide 21: Townsend avalanches
Multiplication starts

## Slide 22: At relatively low voltages Townsend avalanches are well localized in space
For single primary electron:
dne=αne dx
ne=exp αx
For n0 primary electrons
Ne = n0exp αx
“Low “ voltages
High voltages
Gas gain A=exp αx

## Slide 23: How Geiger used both effects(drift and multiplication) in the invented him detector?

## Slide 24: Avalanche development
The process of avalanche multiplication was studied by J. S. Townsend in gases at pressures of several Torr. The discovery was commented by E. Rutherford and H. Geiger:" Suppose that the current between two parallel plates immersed in a gas at low pressure is observed when the air was ionized by X-rays. The current though the gas for small voltages at first increases with the field and then reaches a saturation value, as is ordinary observed in ionized gases at atmospheric pressure. When the field is increase beyond a certain value, however, the current rises rapidly. Townsend has shown that this effect is due to the production of fresh ions in the gas by the collision of the negative ions with the gas molecules. At a later stage, when the electric field approaches the value required to cause a spark, the positive ions also become effective as ionizers but to a much smaller degree that the negative. Under such condition, the small current through the gas due the external ionizing agency may be easily increased several hundreds times. The magnification of the current depends upon the voltage applied and becomes very large just below the sparking value”.

## Slide 25: Geiger counters
Proportional mode
Geiger mode
E=1/r

## Slide 26: In the beginning of the last century , when the electronics just started to be developed, the Geiger mode was the most useful since it provides with a huge output signal

## Slide 27: Quenching of discharge in Geiger mode
( the same principle is now used in
some other detectors: resistive plate
chambers and in SiPM-a compact solid-
stae photomultiplier)

## Slide 28: In Geiger mode the signal was few hundreds volts!
However, in this mode the detector does not record directly particle tracks:
the entire detector volume is under the corona discharge, so any position information is lost

## Slide 29: Why even proportional Geiger detectors do not allow to measure the particle tracks in an easy way?

## Slide 30
The first version of proportional counters has not any imaging capability
(Later one-dimensional single wire counters were developed)

## Slide 31: For this reason for several decades cloud chambers, bubble chambers and spark counters were mostly used in various experiments

## Slide 32: Tracks imaging device
Wilson
~1911

## Slide 33: Images obtained with Wilson chamber

## Slide 34: Home-made cloud chamber (you can make it yourself)

## Slide 35: At the beginning of 30th last century another electron amplifier appear-PMT
PMT is a great device, but was is not a position- sensitive
Zvorykin

## Slide 36: Bubble chamber (1952)
Glaser
most often liquid hydrogen

## Slide 37: Single spark counters (spark and streamer chambers were most widely used as research tools from the 1930s to the 1960s)

## Slide 38
Later there were attempts to introduce electronic readouts
Stacks of spark counters
The most common way of track
detection was the photography
of sparks.

## Slide 39: The main drawbacks of cloud, bubbler and spark chambers- low repetition rate of detection. Moreover, spark chambers could not detect simultaneously many events
In one word: these detectors did not fit growing requirements of high-energy physics experiments and new applications

## Slide 40: MWPC was the first high speed imaging electronic detectors capable recording continuously and simultaneously many particles or photons
G. Charpak

## Slide 41
How MWPC operates?
Equivalent of an array of proportional single –wire counters .
Rate capability ~105 count/sec per wire. 1-D pos. resol. 2.5 mm

## Slide 42
Modern sophisticated MWPC
Position resolution 0.1 - 0.3 mm
Position resolution approaching 0.2 X 0.2 mm

## Slide 43
Photograph of a medium
size MWPC

## Slide 44: MWPCs open the era of modern high efficiency electronic imaging devices
Their use in medical and biological studies
will be reviewed in the second lecture

## Slide 45: At this time it was a revolution in the detection technique
Of course, later solid state
and other competing
devices appeared

## Slide 46: Solid-state micropattern detectors

## Slide 47: Microchannel plates

## Slide 48
Space resolution ~10 µm, time rsolution <100 ps

## Slide 49
ALICE RICH-
each module
40x60cm2
8x8 inch
Vacuum PMT
Vacuum MCP

## Slide 50: Another example of large area gaseous detector-RPC

## Slide 51: RPC
Main advantage - a continuously sensitive device having high time resolution ~1ns

## Slide 52: These detector also can be build with large sensitive areas- in total thousands of m2!
To build them is much easier than MWPCs!

## Slide 53: Microgap RPCs
The main advantage of microgap
RPCs is high time resolution- about
50-70 ps per passing particle and a very
high position resolution-better than 50 μm

## Slide 54: Another recent revolution- micropattern gaseous detectors

## Slide 55: At the end of the last century, or be more precise, at the end of 90-s a new revolution happens in gaseous detectors: micropattern detectors were invented
Their main feature is that they are manufacture by a microelectronic technology and this offers:
● Easy way of manufacturing
● Low cost
● High granularity-high position resolution

## Slide 56
A. Oed

## Slide 57: Micropixel gaseous detectors-something very promising!

## Slide 58: High granularity pixelized readout
The latest tendency: a strip-type of the micropattern detectors readout is replacing by a pixel- type of the readout. This becomes possible due to the fast development of electronics

## Slide 59
GEM and VLSI pixel ASIC @ INFN Pisa
Three ASIC generations of increasing size, reduced pitch & improved functionality
0.18 mm CMOS VLSI
6 layers 0.18 mm CMOS technology
Self-Triggering capabillity
50 * 50 mm pixels (470 pixels/mm2)
Serial analog readout for each pixel
Noise ENC ~ 50 e-/channel
E .Costa et al., Nature 411 (2001) 662
R. Bellazzini et al., NIMA535 (2004) 477
R. Bellazzini et al., NIMA560 (2006) 425
R. Bellazzini et al., NIMA566 (2006) 552
Electron extraction from GEM to pixels

## Slide 60: MICROMEGAS+pixelized readout
Electron from traversing particle drifts towards Micromegas grid and is focused into one of the holes
Thereafter a gas avalanche is induced ending directly at the anode pad of the pixel chip ( no electron extraction from the multiplication structure)
~ -500 V
InGrid
F. F. Hartjes Report at Micropattern Conf.
Crete, June 2009
…
Comparatively low drift field (100 -700 V/mm)
Comparatively low drift field (100 -700 V/mm)
High amplification field (~ 10 kV/mm) to induce gas avalanche
Micromegas holes centred on pads pixel chip

## Slide 61
“InGrid’ Technology and “Driving” Developments
2011: Major Step Forward 
InGrid Production on a wafer level (107 chips)
2005: Single “InGrid” Production
2009: “InGrid” Production on a
3 x 3 Timepix Matrix

## Slide 62
"InGrid" Detector: Single Electron Response and Discharges
Fe55
source
Observe electrons (~220) from an X-ray (5.9 keV) conversion one by one and count them in micro-TPC (6 cm drift)  Study single electron response
1.5 cm
P. Colas, RD51 Collab. Meet.,
Jun.16-17, 2009, WG2 Meeting
Provoke discharges by introducing small amount of Thorium in the Ar gas - Thorium decays to Radon 222 which emits 2 alphas of 6.3 & 6.8 MeV
 Round-shape images of discharges
M. Fransen, RD51 Collab. Meet.,
Oct.13-15, 2008, WG2 Meeting

## Slide 63: This is cutting edge of the technology!
(medical application of these novel detectors, probably , will come soon)

## Slide 64: Liquid detectors and their combination with gaseous detectors-another promising direction (to be discussed later)

## Slide 65: History of gaseous detectors developments and their main designs
RPC

## Slide 66: In the next lectures we will review the main designs of gaseous detectors:
Wire type
Parallel-plate type
Micropattern