## Slide 1
PhD Student: Arianna Elefante
DEVELOPMENT OF INNOVATIVE OPTO-ACOUSTIC SENSORS FOR INDUSTRIAL AND BIOMEDICAL APPLICATIONS
Dipartimento INTERATENEO DI FISICA <<M. MERLIN>>
PhD SCHOOL OF PHYSICS
XXXIII cycle
Second year report

## Slide 2
OUTLINE
2° year PhD Research activities:
- Compact Quartz Enhanced Photoacoustic sensor for the detection of SF6
- Design of an innovative Acoustic Detection Module with two Quarz Tuning Forks
- Study of the stability of a compact QEPAS sensor for the CH4 detection
- Calibration of the CH4 QEPAS sensor with the water vapor concentration
3° year PhD goals
Publications and conference proceedings

## Slide 3
Compact Quartz Enhanced Photoacoustic sensor for the detection of SF6
LASER
DFB QCL
λ= 10.54 μm
LENS
f=50 mm
QTF
L=9.4 mm
T=0.5 mm
S=0.8 mm
Microresonator Tubes
L=12.4 mm
ID=1.51 mm
Acoustic Detection Module
+
+
Motivation: development of a compact QEPAS sensor to be implemented in a test station for the leakage detection in mechatronic systems
1) Detection

## Slide 4
Laser driver + TEC
Main Board
RedPitaya
Power supply
2) Laser Driver and TEC; Electronic for the acquisition and generation of the QEPAS signal
3) Gas line
Flow meter Axetris
Pressure Controller Alicat
ADM
Thomas Pump
Manual valve

## Slide 5
Minimum detection limit:
10 ppb at 8 ms integration time.
Calibration curve
Compact Quartz Enhanced Photoacoustic sensor for the detection of SF6
Hitran simulation of the absorption cross-section for pure SF6 at 80 Torr
QEPAS spectral scan measured for 20 ppm of SF6 at 80 Torr

## Slide 6
Dual QTF ADM
Motivation: analysis of mixtures with a wide dynamic range in concentration from few ppm to the percentage range
Non-linear effects:
Molecular relaxation dynamics
Speed of sound
S1 : bare QTF
S2: QTF + AmR tubes
Sensitive measurements of low concentration
High concentration measurement

## Slide 7
Dual QTF ADM
228 mm
S1
S2
Adapter for the alignment of S1
Alignment of S2
5 mm
55.6 x 50.6 x 32.6 mm

## Slide 8
Study of the stability of a compact QEPAS sensor for the CH4 detection
Permselect
Ultimate Detection Limit:
57 ppb @ Tint= 8 ms
2 ppb @ Tint= 6 s

## Slide 9
Calibration of the CH4 QEPAS sensor with the water vapor concentration
Motivation: Normalize the CH4 QEPAS signal with respect to the absolute humidity
CH4 QEPAS signal
H2O Absolute Humidity TSP
H2O QEPAS signal

## Slide 10
Calibration of the CH4 QEPAS sensor with the water vapor concentration
CH4 (mV)=0.00044 (mV/ppm) * H2O AH (ppm) + 2. 09733 mV
R2=0. 982
CH4 QEPAS (mV)

## Slide 11
3° year PhD goals
Implementation of the SF6 sensor in the test station at Masmec
Assembling and testing of the QEPAS sensor with the dual QTF ADM for the detection of CH4 in a wide concentration range

## Slide 12
Publications and conference proceedings
Publications
M. Giglio, A. Zifarelli, A. Sampaolo, G. Menduni, A. Elefante, R. ,Blanchard, C. fluegl,M.F. Witinski, D. Vakhshoori, H. Wu, V.M.N. Passaro, P. Patimisco, F.K. Tittel, L. Dong, V. Spagnolo, “Broadbanddetection of methane and nitrous oxide using a distributed feedbackquantum cascade laser array and quartz-enhanced photoacousticsensing”, Photoacoustics 17, 100159 (2020).
A. Elefante, M. Giglio, A. Sampaolo, G. Menduni, P. Patimisco, V. M.N. Passaro, H. Wu, H. Rossmadl, V. Mackowiak, A. Cable, F.K. Tittel, L. Dong and V. Spagnolo, “Dual-Gas Quartz-Enhanced Photoacoustic Sensor for Simultaneous Detection of Methane/Nitrous Oxide and Water Vapor”, Anal. Chem. 91, 12866-12873 (2019).
Conference Proceedings
Fiber-coupled quartz-enhanced photoacoustic sensor for methane and ethane trace detection, Fabrizio Sgobba, Univ. degli Studi di Bari Aldo Moro (Italy); Giansergio Menduni, Angelo Sampaolo, Pietro Patimisco, Marilena Giglio, Politecnico di Bari (Italy); Arianna Elefante, Univ. degli Studi di Bari Aldo Moro (Italy); Vittorio M. N. Passaro, Politecnico di Bari (Italy); Frank K. Tittel, Rice Univ. (USA); Vincenzo Spagnolo, Politecnico di Bari (Italy)
Measurement of non-radiative gas molecules relaxation rates by using quartz-enhanced photoacoustic spectroscopy, Stefano dello Russo, Univ. degli Studi di Bari Aldo Moro (Italy); Pietro Patimisco, Angelo Sampaolo, Marilena Giglio, Giansergio Menduni, Politecnico di Bari (Italy); Arianna Elefante, Univ. degli Studi di Bari Aldo Moro (Italy); Vittorio M. N. Passaro, Politecnico di Bari (Italy); Frank K. Tittel, Rice Univ. (USA); Vincenzo Spagnolo, Politecnico di Bari (Italy).
Partial least squares regression as novel tool for gas mixtures analysis in quartz-enhanced photoacoustic spectroscopy, Andrea Zifarelli, Univ. degli Studi di Bari Aldo Moro (Italy); Pietro Patimisco, Angelo Sampaolo, Marilena Giglio, Giansergio Menduni, Politecnico di Bari (Italy); Arianna Elefante, Univ. degli Studi di Bari Aldo Moro (Italy); Vittorio M. N. Passaro, Politecnico di Bari (Italy); Frank K. Tittel, Rice Univ. (USA); Vincenzo Spagnolo, Politecnico di Bari (Italy)